## Slide 1: Next Generation Sequencing – An Overview
Olga Vinnere Pettersson, PhD
National Genomics Infrastructure hosted by ScilifeLab,
Uppsala Node (UGC)
Version 6.2

## Slide 2: Outline:
A bit of history
NGS technologies & sample prep
NGS applications
National Genomics Infrastructure – Sweden
www.robustpm.com

## Slide 3: HISTORY

## Slide 4: What is sequencing?
https://figures.boundless-cdn.com

## Slide 5: Once upon a time…
Fredrik Sanger and Alan Coulson
Chain Termination Sequencing (1977)
Nobel prize 1980
Principle:
SYNTHESIS of DNA is randomly TERMINATED at different points
Separation of fragments that are 1 nucleotide different in size
Lack of OH-group at 3’ position of deoxyribose
!
1 molecule sequenced at a time = 1 read
Capillary sequencer: 384 reads per run

## Slide 6: 2006 REVOLUTION
Thousands of molecules sequenced in parallel
1 mln reads sequenced per run
Roche 454 GS FLX

## Slide 7
Massively parallel sequencing (454, Illumina, Life Tech)
Human genome
James Watsons genome
DNA sequencing revolution - Sweden
Center for Metagenomic
Sequence Analysis (KAW)
Science for Life
Laboratory (SciLifeLab)
National Genomics Infrastructure
(NGI)
Notes: Tight schedule, Lunch (wrap) at level 2, Dinner at level 2, speakers hand in USB, Thanks to Anders Andersson

## Slide 8
Workload at NGI – Sweden 2010-2015
• 30% increase in number of NGS samples from previous year (36 000)

## Slide 9
Technologies

## Slide 10: Making an NGS library
DNA QC – paramount importance
Sharing & size selection
Ligation of sequencing adaptors, technology specific
Amplification

## Slide 11: NGS technologies
RIP technologies: Helicos, Polonator, SOLiD, 454 etc.
In development: Tunneling currents, nanopores, etc.
Company | Platform | Amplification | Sequencing method
Roche | 454 (until 2016) | emPCR | Pyrosequencing
Illumina | HiSeq, MiSeq NextSeq, X10 | Bridge PCR | Synthesis
LifeTechnologies(Thermo Fisher) | Ion Torrent, Ion Proton, S5 | emPCR | Synthesis (pH)
Pacific Biosciences | RSII | None | Synthesis (SMRT)
Complete genomics | Nanoballs | None | Ligation
Oxford Nanopore* | MinION GridION | None | Flow

## Slide 12: Differences between platforms
Technology: chemistry + signal detection
Run times vary from hours to days
Production range from Mb to Gb
Read length from <100 bp to > 20 Kbp
Accuracy per base from 0.1% to 15%
Cost per base

## Slide 13: Illumina
Main applications
Whole genome, exome and targeted reseq
Transcriptome analyses
Methylome and ChiPSeq
Rapid targeted resequencing (MiSeq)
Human genome seq (Xten)
Instrument | Yield and run time | Read Length | Error rate | Error type
HiSeq2500 | 120 Gb – 600 Gb 27h or standard run | 100x100 (250x250) | 0.1% | Subst
MiSeq | 540 Mb – 15 Gb (4 – 48 hours) | Up to 350x350 | 0.1% | Subst
HiSeqXten | 800 Gb - 1.8 Tb (3 days) | 150x150 | “ | “

## Slide 14: Illumina sequencing - bridge amplification

## Slide 15: Ion Torrent
Main applications
Microbial and metagenomic sequencing
Targeted re-sequencing (gene panels)
Clinical sequencing
Chip | Yield - run time | Read Length
314, 316, 318 (PGM) | 0.1 – 1 Gb Gb, 3 hrs | 200 – 400 bp
P-I (Proton) | 10 Gb 4 hrs | 200 bp
520, 530, 540 (S5) | 1 Gb – 10 Gb 3 hrs | 400 bp (except 540)

## Slide 16: Ion Torrent - H+ ion-sensitive field effect transistors

## Slide 17
314 316 318
PI
250 000 4 mln 9 mln
400 bp 400 bp 400 bp
100 Mb 500 Mb 1 Gb
Ion PGM
Ion Proton
520 530 540
8 mln 15-20 mln 90 mln
400 bp 400 bp 200 bp
1 Gb 5 Gb 10 Gb
90 mln
200 bp
10-18 Gb
Ion S5XL

## Slide 18: PacBio SMRT-technology
Instrument | Yield and run time | Read Length | Error rate | Error type
RS II Sequel (R&D) | 250 Mb – 1.3 Gb /30 - 360 min SMRTCell | 250 bp – 30 000 bp (70 000 bp) | 15% (on a single passage!) | Insertions, random
Single-Molecule, Real-Time DNA sequencing

## Slide 19: PacBio SMRT - technology
Single Molecule Real Time

## Slide 21
Oxford Nanopore MinION
Reads up to 100k
1D and 2D reads
15-40% error rate
Life time 5 days

## Slide 22: Main types of equipment
Illumina HiSeq
Illumina Xten
Illumina MiSeq
Short paired reads
HIGH throughput
Ion Torrent PGM
Ion Proton
Ion S5 XL
Short single-end reads
FAST throughput
PacBio RSII
Ultra-long reads
FAST throughput

## Slide 23
Applications

## Slide 24: NGS/MPS applications
Whole genome sequencing:
  De novo sequencing
  Re-sequencing
Transcriptome sequencing:
  mRNA-seq
  miRNA
  Isoform discovery
Target re-sequencing
  Exome
  Large portions of a genome
  Gene panels
  Amplicons

## Slide 25: De novo sequencing
Used to create a reference genome without previous reference

## Slide 26: Transcriptome sequencing (RNA-seq)
TOTAL RNA
mRNA
Dif.ex.
Annotation
miRNA
Non-codingRNA
Splice isoforms
Transcriptional regulation

## Slide 27: mRNA: rRNA depletion vs polyA selection
Method | Pros | Cons | Recommended
rRNA depletion | Captures on-going transcription Picks up non-coding RNA | Does not get rid of all rRNA Messy Dif.Ex. profile | 20-40 mln reads (single or PE)
polyA selection | Gives a clean Dif.Ex. profile | Does not pick non-coding RNA | 5-20 mln reads
Alternative for human RNA-seq:
AmpliSeq Human Transcriptome panel:
faster, cheaper, works fine with FFPE
input: 50 ng total RNA
dif.ex. ONLY

## Slide 28: RNA-seq experimental setup
mRNA only: any kit
mRNA and miRNA: only specialized kits
Always use DNase!
RIN value above 8.
CONTROL vs experimental conditions
Biological replicates: 4 strongly recommended

## Slide 29: Amplicon sequencing
Used a lot in metagenomics
Community analysis
  rRNA genes & spacers (16S, ITS)
  Functional genes
Genotyping by sequencing

## Slide 30: Amplicon sequencing
Example 2: several sizes,
fractionation is needed
=> we HAVE to make several libraries
Example 3: broad peak;
size selection is needed
FOR ANY NGS TECHNOLOGY
Size difference among fragments must not
exceed 80 bp (or 20% in length)
Reason – preferential amplification of short fragments
Example 1: tight peak, OK
SIZE MATTERS…

## Slide 31: Size-related bias in amplicon-seq
Courtesy Mikael Brandström Durling, Forest Mycology and Pathology, SLU

## Slide 32: When you sequence an amplicon…
On MiSeq
FW read
RW read
FW read
On Ion

## Slide 33: Main types of equipment & applications
Illumina HiSeq
Illumina Xten
Illumina MiSeq
Short paired reads
HIGH throughput
Human WGS
mRNA and miRNA
De novo transcriptome
Exome
ChIP-seq
Short amplicons
Methylation
Ion Torrent PGM
Ion Proton
Ion S5 XL
Short single-end reads
FAST throughput
mRNA and miRNA
Exome
ChIP-seq
Short amplicons
Gene panels
Clinical samples
PacBio RSII
Ultra-long reads
FAST throughput
Long amplicons
Re-sequencing
De novo sequencing
Novel isoform discovery
Fusion transcript analysis
Haplotype phasing
Clinical samples

## Slide 34: What is The BEST?

## Slide 35: SAMPLE QUALITY REQUIREMENTS

## Slide 36: Sample prep: take home message
PCR-quality sample and
NGS-quality sample
are two completely different things

## Slide 37: What do absorption ratios tell us?
Pure DNA 260/280: 1.8 – 2.0
< 1.8:
Too little DNA compared to other components of the solution; presence of organic contaminants: proteins and phenol; glycogen - absorb at 280 nm.
> 2.0:
High share of RNA.
Pure DNA 260/230: 2.0 – 2.2
<2.0:
Salt contamination, humic acids, peptides, aromatic compounds, polyphenols, urea, guanidine, thiocyanates (latter three are common kit components) – absorb at 230 nm.
>2.2:
High share of RNA, very high share of phenol, high turbidity, dirty instrument, wrong blank.
Photometrically active contaminants:
phenol, polyphenols, EDTA, thiocyanate, protein,
RNA, nucleotides (fragments below 5 bp)

## Slide 38: How to make a correct measurement
Thaw DNA completely
Mix gently (never vortex!)
Put the sample on a thermoblock: 37°C, 15-30 min
Mix gently
Dilute 1:100 (if HMW)
Mix gently
Make a measurement with an appropriate blank
NANODROP is Bad. Point.
Use Qubit, or PicoGreen.
Low concentration
High concentration
DNA
solution

## Slide 39: Let’s get philosophical

## Slide 40: Since the beginning of Genomics:
First genome: virus  X 174 - 5 368 bp (1977)
First organism: Haemophilus influenzae - 1.5 Mb (1995)
First eukaryote: Saccharomyces cerevisiae - 12.4 Mb (1996)
First multicellular organism: Cenorhabditis elegans - 100 MB (1998-2002)
First plant: Arabidopsis thaliana - 157 Mb (2000)

## Slide 41
… prices go down
Human genome project, 2007
  Genome of Craig Wenter costs 70 mln $
    Sanger’s sequencing
  Genome of James Watson costs 2 mln $
    454 pyrosequencing
  Ultimate goal: 1000 $ / individual
    Almost there! (1200 $)

## Slide 42: … paradigm change
  From single genes to complete genomes
  From single transcripts to whole transcriptomes
  From single organisms to complex metagenomic pools
  From model organisms to the species you are studying
  Personal genome = personalized medicine

## Slide 43
IF 2.9
IF 31.6
… scientific value diminishes

## Slide 44: Main challenge - DATA ANALYSIS and DATA STORAGE
http://finchtalk.geospiza.com
!
=> More bioinformaticians to people!
$
Sequencing
Data analysis

## Slide 45: NGI-portal

## Slide 46: Good reading:
http://journals.plos.org/plosbiology/article?id=10.1371/journal.pbio.1001744

## Slide 48: National Genomics Infrastructure
Mid 2010
SciLifeLab, Stockholm
SciLifeLab, Uppsala
Uppmax, Uppsala

## Slide 49
10 Illumina HiSeq Xten
17 Illumina HiSeq 2000/2500
3 Illumina MiSeq
1 Illumina NextSeq
2 Life Technologies Ion Torrent
6 Life Technologies Ion Proton
2 Pacific Biosciences RSII
2 Sanger ABI3730
1 Argus Whole Genome Map. Syst.
2 Oxford Nanopore MinIon
NGI-SciLifeLab is one of the most well-equipped
NGS sites in Europe
Notes: NGI is one of the most well equipped NGS site in Europe. We use different technologies which makes it possible to find most appropriate solution for different type of projects
49

## Slide 50
https://ngisweden.scilifelab.se/

## Slide 51
What happens then?
NGI Project coordinators meet twice a week via Skype
Project is then assigned to a certain node and a coordinator contacts the PI
Project distribution is based on:
Wish of PI
Type of sequencing technology
Type of application
Queue at technology platforms
Ulrika Ellenor
Liljedahl Devine
SNP&SEQ, Uppsala node
Mattias Beata
Ormestad Werne Solenstam
Stockholm Node
Olga
Vinnere Pettersson
UGC, Uppsala Node

## Slide 52
Projects at CMS
3. Access to genomics platform
Project meeting
What we can help you with:
Design your experiment based on the scientific question.
Chose the best suited application for your project.
Find the most optimal sequencing setup.
Answer all questions about our technologies and applications, as well as bioinformatics.
In special cases, we can give extra-support with bioinformatics analysis – development of novel methods and applications

## Slide 53
Need help with data analysis?
NBIS.se

## Slide 54
QUESTIONS?